## Slide 1
A Carbon-Neutral,
Emission-Free,
Particle Accelerator
for Ion Beam Analysis
Dan Faircloth, Olli Tarvainen
ISIS Low Energy Beams Group
UKRI STFC Rutherford Appleton Laboratory
Taneli Kalvas, Ville Toivanen
Accelerator Laboratory
University of Jyväskylä, Finland

## Slide 2
Particle accelerators traditionally use a lot of power…
ISIS uses approximately 11 MW
to produce a 160 kW proton beam
(which is then dumped into the targets!)
The only solution is to build a communal heating system that can be switched off when ISIS is running

## Slide 3
…or require large amounts of SF6 (for high voltage insulation)
…or use a design that eliminates the need for insulating gas entirely:
SF6 is the most potent greenhouse gas:
1 kg SF6 “lost” = 23 tonnes of CO2
stays in the atmosphere for 3,200 years
A possible solution is to modify the system to use a modern alternative
e.g.
C4-FN (fluoronitrile) mixture
(O2 / CO2 / N2)
GE and Hitachi

## Slide 4
RF ion source or
duoplasmatron
e.g. He+
5-8 bar SF6
sample
detector
high voltage terminal
Typical IBA Acceleration Schemes and Sources
1-20 MV
“single ended”
very compact and hard to access ion source
positive ion beam

## Slide 5
sample
detector
stripper
(gas cell or foil)
high voltage terminal
Typical IBA Acceleration Schemes and Sources
A zoo of difficult to use and unstable negative ion sources:
Caesium sputter
Charge exchange
Volume
e.g. Cl- , H-
“tandem”
5-8 bar SF6
positive ion beam
negative ion beam
double your energy for free
1-20 MV
negative
ion source

## Slide 6
sample
detector
high voltage terminal
high charge state
positive ions
Emission-Free, Zero-Carbon, IBA Concept:
500 kV
air insulated (SF6 free)
high voltage design
e.g. Xe27+
ECR ion source
q/m dipole
use high charge state ions = much lower accelerating voltage required
This approach obviously rules out using light ions, however material modification and
ToF ERDA and are still possible

## Slide 7
Time-of-Flight Elastic Recoil Detection Analysis (ToF-ERDA)
Best suited for light elements including H, Li, C, N and O.
These elements cannot be detected using photon-based (x-ray) methods.
Hydrogen is “invisible” to Rutherford Backscattering Spectrometry (RBS), which makes ToF-ERDA unique.
Simple principle:
Bombard the sample with heavy ions
Measure the velocity and energy of the recoiling particles
Use known energy loss in material to obtain depth profile
Lighter than beam particles are ejected, scattered beam
provides information on heavier elements.
A method to measure the elemental depth profile in materials

## Slide 8
Time-of-Flight Elastic Recoil Detection Analysis (ToF-ERDA)
Example: metal tri-layer on silicon-oxide substrate
Useful tool for developing thin-film/surface technologies:
Superconducting RF
Electronic devices
Sensors
Functional surface coatings
Catalytic surface chemistry
Energy applications
Battery technology
…

## Slide 9
Typical ToF-ERDA ion beam requirements
Beam | Energy
35/37Cl | 3 - 6 MeV
79/81Br | 8 - 11 MeV
127I | 9 - 16 MeV
The required particle flux at the sample is 1-10 pnA
Lower limit: Reasonable measurement time, typically less than an hour / sample.
Upper limit: The maximum count rate of the energy detector.

## Slide 10
… to limit the platform voltage to 500 kV
Beam | Energy
35/37Cl | 3 - 6 MeV
79/81Br | 8 - 11 MeV
127I | 9 - 16 MeV
Ion | Charge state required | Energy
40Ar | 6+ … 12+ | 3 - 6 MeV
84Kr | 16+ … 22+ | 8 - 11 MeV
136Xe | 18+ … 32+ | 9 - 16 MeV
1-10 pnA of
these charge states
we need
We need an Electron Cyclotron Resonance (ECR) ion source
it needs to be a permanent magnet based ion source

## Slide 11
Permanent magnet ECR ion sources
SUPERNANOGAN
DECRIS-PM

## Slide 12
1~5：PM rings
6, 7：soft iron rings
8~11：soft iron plates
12~14：auxiliary elements,
15：hexapole
16：coil
525 kg of PM material
Dubna ECR Ion Source

## Slide 13
Commercial permanent magnet ECRIS
approx. £1M

## Slide 14
Permanent magnet CUBE-ECRIS
Novel quadrupole minimum-B structure
160 kg of PM material

## Slide 15
Permanent magnet CUBE-ECRIS
CUBE-ECRIS
Electrostatic quadrupole doublet
m/q-analysis magnet
(4) Faraday cup
(5) Emittance scanner

## Slide 16
Permanent magnet CUBE-ECRIS
Argon charge state distribution with 15-nitrogen buffer gas to measure Ar12+ without overlapping buffer gas ions
Charge state | Current [A] | Flux [pnA]
5+ | 25 | 5000
6+ | 18 | 3000
7+ | 15 | 2140
8+ | - | -
9+ | 12 | 1330
10+ | 5.8 | 580
11+ | 1.5 | 140
12+ | 0.4 | 33
12 kV source potential
TOF ERDA suitable
@ 500 kV

## Slide 17
Permanent magnet CUBE-ECRIS
Krypton (enriched 84-isotope) charge state distribution
Charge state | Current [A] | Flux [pnA]
17+ | 1.25 | 100
18+ | 0.73 | 40
19+ | 0.23 | 12
TOF ERDA suitable
@ 500 kV
(9.5 MeV)

## Slide 18
Permanent magnet CUBE-ECRIS
Xenon (enriched 136-isotope) charge state distribution
Charge state | Current [A] | Flux [pnA]
18+ | 4.9 | 270
19+ | 4.55 | 240
20+ | 2.7 | 135
21+ | 1.5 | 70
22+ | - | -
23+ | 0.5* | 22*
24+ | 0.2 | 8
*Enriched 131-isotope
TOF ERDA suitable
@ 500 kV
(12 MeV)

## Slide 19
Permanent magnet CUBE-ECRIS
One possibility to increase the beam currents of high charge state Kr and Xe beams is to increase the microwave power.
The high charge state currents are power limited as shown here for argon.

## Slide 20
Permanent magnet CUBE-ECRIS
One possibility to increase the beam currents of high charge state Kr and Xe beams is to increase the microwave power.
The high charge state currents are power limited as shown here for krypton (enriched 84Kr isotope).

## Slide 21
Permanent magnet CUBE-ECRIS
One possibility to increase the beam currents of high charge state Kr and Xe beams is to increase the microwave power.
The high charge state currents are power limited as shown here for xenon (enriched 131Xe isotope).

## Slide 22
Permanent magnet CUBE-ECRIS
CUBE-ECRIS
Electrostatic quadrupole doublet
m/q-analysis magnet
(4) Faraday cup
(5) Emittance scanner

## Slide 23
Adjustable field permanent magnet dipole for charge state selection
funded by:
Physics design and simulation
Engineering design and assembly
Field measurement and validation

## Slide 24
Adjustable field permanent magnet dipole for charge state selection
Perfect match between simulation and
measured field with varying tuner position
Excelllent field uniformity radially across the magnet pole,
B/B < 5x10^-4 (the specification is 1x10-3)
Other potential applications of the magnet:
accelerator based nuclear physics
accelerator mass spectrometry
isotope separation

## Slide 25
First beam transport tests with the PM dipole – May 2024
PM dipole magnet

## Slide 26
First beam transport tests with the PM dipole – May 2024
Argon beam currents in A
Argon beam currents in pnA

## Slide 27
First beam transport tests with the PM dipole – May 2024
Argon beam currents are lower than with the electromagnet
This is due to lack of optimisation:
No double frequency heating
Higher than normal drain current
Beam losses due to severe extraction misalignment
Beam transport simulations of the whole LEBT to confirm
the effect on total transmission
Puller electrode offset and rotation wrt the extraction slit

## Slide 28
First beam transport tests with the PM dipole – May 2024
Puller electrode offset and rotation wrt the extraction slit
Asymmetric beam losses in the puller electrode
and poor quality beam into the LEBT

## Slide 29
First beam transport tests with the PM dipole – May 2024
Beam burn mark matching the simulation with puller offset
Beam induced sputtering
coating of the puller feedthrough
failure ending the experiment
Requires correcting the misalignment and
changing the puller feedthrough location

## Slide 30
Next steps
Continue optimisation of front end
Develop the rest of the system

## Slide 31
sample
detector
e.g. Xe27+
ECR ion source
high charge state
positive ions
q/m dipole
Emission-Free, Zero-Carbon, IBA Concept – Required Components
Battery power on platform
Commercially available post insulators
Commercially available accelerating column
Commercially available experiment station
Safety, control and integration
500 kV
DC

## Slide 32
5 x 100 kV NEC air-insulated acceleration tubes received (100 W leakage power)
500 kV NEC

## Slide 33
Beam transport simulations
Same overall transport efficiency for higher charge states - 5 MeV Ar10+
Taneli Kalvas, University of Jyvaskyla
PM dipole
500 kV acceleration tube (NEC)
Ar10+
Ar11+

## Slide 34
Courtesy of Mikko Laitinen, University of Jyvaskyla
Modern ToF-ERDA tool:
state-of-the-art hydrogen detection
Efficiency
automation options
University of Jyvaskyla

## Slide 35
Carbon Neutral!
24x7 base load: 2 kW
+ when operating: 1.5 kW
Operating 8 hours/day
5 days/week
= 20,000 kWh per annum
Based on UK average:
11 m x 11 m of solar panels required
(approximately the building footprint)

## Slide 36
Summary
Power consumption low enough to use 100% local solar
Zero emission - SF6 eliminated
Reliable, flexible and stable ion source
Reduced maintenance
Turn-key/push button system
A critical tool for developing thin-film/surface technologies

## Slide 37
Thankyou Questions?